## Slide 1
Hometown: Kayenta, Arizona
(Navajo: Tó Dínéeshzheeʼ)
Alastair Big Luna (AL)
Alastair Big Luna (AL)

## Slide 2: Optical Fiber Temperature and Stress Measurements

## Slide 3
Section 1
Discussion on how optical fibers work
Discussion on Rayleigh backscatter
Discussion on OFDR
Discussion on the temperature and strain algorithm
Section 2
Characteristics of the LUNA ODiSI-B and temperature data
Research on using strain data on pressurized vessels to characterize internal pressure
Presentation Outline

## Slide 4: Section 1: Background
Optic fibers, Rayleigh scattering, OFDR, Temperature & Strain Algothrim

## Slide 5
Optical fibers may be used for temperature, strain, pressure.
The spatial resolution is very high. Choice of having 1.25 [mm] at 24 [Hz] or 5 [mm] at 100 [Hz] of spacing between each data point.
  For example, in a one-meter sensor, you can have as much 800 points of data for each time step. At 24 [Hz], you have 19,200 points of data for each second.
For this reason, fiber optic sensors are candidates for research, big data collection for machine learning algorithms, and health monitoring of critical components in power plants or facilities.
Why use fiber optic sensors
Figure: Fiber-optic temperature and strain sensor – 1 [m] length

## Slide 6: Required components for temperature and strain measurements
Single mode fibers
Tunable laser and reflectometer
Temperature and Strain algorithm

## Slide 7
Snell’s law
The critical angle
The bigger the difference between the indices of refraction in the two optical materials, the larger the critical angle.
The fiber is design to be a light guide.
Optical Fiber Theory – Total Internal Reflection
Figure: Illustration of Total Internal Reflection [1]
[1] https://images.fineartamerica.com/images-medium-large-5/refraction-and-total-internal-reflection-russell-kightley.jpg

## Slide 8
Outside Vapor Deposition (OVD) process is commonly used in the manufacturing of optical fibers.
3 steps: Laydown, consolidation, and draw.
  Laydown process - a soot preform is made from ultra pure vapors as they react in the flame to form fine soot particles of silica or Germania.
  Consolidation process – the bait rod is removed and the preform is placed in a furnace. The preform is sintered into a solid glass.
  Draw process – the preform is placed in a draw tower and drawn into one continuous strand of silica fiber.
Silica Optical Fiber – Manufacturing process
Figure: Illustration of Laydown process [2]
[2] P. Aithal and H. Ravindra, Textbook of Engineering Physics, 1st ed. Acme Learning Private Limited, 2021.

## Slide 9
First proposed for measuring losses inside single-mode optical fiber system. [3]
Relies on a scan of the Rayleigh backscatter amplitude.
A frequency scan is made by changing the wavelength of the laser and a detector measures the backscatter amplitude
Optical Frequency Domain Reflectometry
Figure 5: Fiber-optic temperature and strain sensor
[3]W. Eickhoff and R. Ulrich, "Optical frequency domain reflectometry in single‐mode fiber", Applied Physics Letters, vol. 39, no. 9, pp. 693-695, 1981. Available: 10.1063/1.92872

## Slide 10
First proposed for measuring losses inside single-mode optical fiber system. [3]
Relies on a scan of the Rayleigh backscatter amplitude.
A frequency scan is made by changing the wavelength of the laser and a detector measures the backscatter amplitude
Optical Frequency Domain Reflectometry
Figure 5: Fiber-optic temperature and strain sensor
[3]W. Eickhoff and R. Ulrich, "Optical frequency domain reflectometry in single‐mode fiber", Applied Physics Letters, vol. 39, no. 9, pp. 693-695, 1981. Available: 10.1063/1.92872

## Slide 11: Optical Frequency Domain Reflectometry (OFDR)-Algorithm
[2]M. Froggatt and J. Moore, "High-spatial-resolution distributed strain measurement in optical fiber with Rayleigh scatter", Applied Optics, vol. 37, no. 10, p. 1735, 1998. Available: 10.1364/ao.37.001735 [Accessed 30 July 2020].
I
ω

## Slide 12: Optical Frequency Domain Reflectometry (OFDR)-Algorithm
[2]M. Froggatt and J. Moore, "High-spatial-resolution distributed strain measurement in optical fiber with Rayleigh scatter", Applied Optics, vol. 37, no. 10, p. 1735, 1998. Available: 10.1364/ao.37.001735 [Accessed 30 July 2020].
FFT
I
ω
X
I
xm

## Slide 13: Optical Frequency Domain Reflectometry (OFDR)-Algorithm
[2]M. Froggatt and J. Moore, "High-spatial-resolution distributed strain measurement in optical fiber with Rayleigh scatter", Applied Optics, vol. 37, no. 10, p. 1735, 1998. Available: 10.1364/ao.37.001735 [Accessed 30 July 2020].
FFT
I
ω
X
I
Gage length
xm

## Slide 14: Optical Frequency Domain Reflectometry (OFDR)-Algorithm
[2]M. Froggatt and J. Moore, "High-spatial-resolution distributed strain measurement in optical fiber with Rayleigh scatter", Applied Optics, vol. 37, no. 10, p. 1735, 1998. Available: 10.1364/ao.37.001735 [Accessed 30 July 2020].
FFT
I
ω
FFT
X
I
Gage length
Reference
ω
I
xm

## Slide 15: Optical Frequency Domain Reflectometry (OFDR)-Algorithm
[2]M. Froggatt and J. Moore, "High-spatial-resolution distributed strain measurement in optical fiber with Rayleigh scatter", Applied Optics, vol. 37, no. 10, p. 1735, 1998. Available: 10.1364/ao.37.001735 [Accessed 30 July 2020].
FFT
I
ω
FFT
X
I
Gage length
Reference
Sample
ω
ω
I
I
xm

## Slide 16: Optical Frequency Domain Reflectometry (OFDR)-Algorithm
[2]M. Froggatt and J. Moore, "High-spatial-resolution distributed strain measurement in optical fiber with Rayleigh scatter", Applied Optics, vol. 37, no. 10, p. 1735, 1998. Available: 10.1364/ao.37.001735 [Accessed 30 July 2020].
FFT
I
ω
FFT
X
I
Gage length
Reference
Sample
Cross Correlated
ω
ω
I
I
Δω
Δωm
xm
I

## Slide 17: Optical Frequency Domain Reflectometry (OFDR)-Algorithm
[2]M. Froggatt and J. Moore, "High-spatial-resolution distributed strain measurement in optical fiber with Rayleigh scatter", Applied Optics, vol. 37, no. 10, p. 1735, 1998. Available: 10.1364/ao.37.001735 [Accessed 30 July 2020].
FFT
I
ω
FFT
X
I
Gage length
Reference
Sample
Cross Correlated
Peak Correlated to a specific location
ω
ω
I
I
Δω
Δωm
xm
xm
Δωm
X
Δω
I

## Slide 18: Optical Frequency Domain Reflectometry (OFDR)-Algorithm
[2]M. Froggatt and J. Moore, "High-spatial-resolution distributed strain measurement in optical fiber with Rayleigh scatter", Applied Optics, vol. 37, no. 10, p. 1735, 1998. Available: 10.1364/ao.37.001735 [Accessed 30 July 2020].
FFT
I
ω
FFT
X
I
Gage length
Reference
Sample
Cross Correlated
Peak Correlated to a specific location
ω
ω
I
I
Δω
Δωm
xm
xm
Δωm
X
Δω
I

## Slide 19: Optical Frequency Domain Reflectometry (OFDR)-Algorithm
[2]M. Froggatt and J. Moore, "High-spatial-resolution distributed strain measurement in optical fiber with Rayleigh scatter", Applied Optics, vol. 37, no. 10, p. 1735, 1998. Available: 10.1364/ao.37.001735 [Accessed 30 July 2020].
FFT
I
ω
FFT
X
I
Gage length
Reference
Sample
Cross Correlated
Peak Correlated to a specific location
ω
ω
I
I
Δω
Δωm
xm
xm
Δωm
X
Δω
I

## Slide 20: Optical Frequency Domain Reflectometry (OFDR)-Algorithm
[2]M. Froggatt and J. Moore, "High-spatial-resolution distributed strain measurement in optical fiber with Rayleigh scatter", Applied Optics, vol. 37, no. 10, p. 1735, 1998. Available: 10.1364/ao.37.001735 [Accessed 30 July 2020].
FFT
I
ω
FFT
X
I
Gage length
Reference
Sample
Cross Correlated
Peak Correlated to a specific location
ω
ω
I
I
Δω
Δωm
xm
xm
Δωm
X
Δω
I

## Slide 21: Optical Frequency Domain Reflectometry (OFDR)-Algorithm
[2]M. Froggatt and J. Moore, "High-spatial-resolution distributed strain measurement in optical fiber with Rayleigh scatter", Applied Optics, vol. 37, no. 10, p. 1735, 1998. Available: 10.1364/ao.37.001735 [Accessed 30 July 2020].
FFT
I
ω
FFT
X
I
Gage length
Reference
Sample
Cross Correlated
Peak Correlated to a specific location
ω
ω
I
I
Δω
Δωm
xm
xm
Δωm
X
Δω
I

## Slide 22: Optical Frequency Domain Reflectometry (OFDR)-Algorithm
[2]M. Froggatt and J. Moore, "High-spatial-resolution distributed strain measurement in optical fiber with Rayleigh scatter", Applied Optics, vol. 37, no. 10, p. 1735, 1998. Available: 10.1364/ao.37.001735 [Accessed 30 July 2020].
FFT
I
ω
FFT
X
I
Gage length
Reference
Sample
Cross Correlated
Peak Correlated to a specific location
ω
ω
I
I
Δω
Δωm
xm
xm
Δωm
X
Δω
I

## Slide 23: Section 2 Distributed Temperature Sensing
LUNA ODiSI-B and ODiSI-6000

## Slide 24: Furnace Setup
Temperature fiber sensors

## Slide 25
High Temperature Characteristics of the LUNA ODiSI-B and SMF

## Slide 26
High Temperature Characteristics of the LUNA ODiSI-B and SMF

## Slide 27
SMF modifications to withstand High Temperature Environments
Test Parameters | 750 [C] - HE | 750 [C] - BN
Interrogator System | Luna ODiSI-B | Luna ODiSI-B
Type of Optical Fiber | Single mode (SMF-28e) | Single mode (SMF-28e)
Heater Type | Tube Furnace | Tube Furnace
Anneal Type | 800 [C] for 2 [hrs] | 800 [C] for 2 [hrs]
Typical Experimental Timeframes | Until failure | Until failure
Termination End Types | ~30 degree polish | ~30 degree polish
Backfilled Gas Type | Helium | Air with Boron Nitride coating
Optical Fiber Coating | Mechanically Removed | Mechanically Removed

## Slide 28: Optical fiber temperature sensors survivability test in 750 [C] furnace
Plot shows the signal loss for the initial keys of Helium fiber (magenta) and BN fiber (blue). Red line shows 50% signal loss threshold. Helium and BN exceeded the 50% line at 18.6 [hr] and 42.4 [hr], respectively.
Corning SMF-28 Ultra can survive 750 [C] temperatures for 1780 hours (10.6 weeks) when protected with Helium gas cover or Boron Nitride coating
If we consider a 50% signal threshold for both He and BN fibers as a metric, then this test shows BN helps the fiber maintain signal more than the Helium backfilled gas by about 24 hours.
31 reference keys were made for both the Helium fiber and BN fiber. Creating numerous reference keys is not ideal to maintain signal, and therefore, we are searching for fibers that have enhanced Rayleigh scatter signals.
18.6 [hrs] – He
42.4 [hrs] – BN

## Slide 29
UW-Madison’s high pressure testing loop can simulate PWR conditions using single or multiple heating rods.
PWR conditions can reach temperatures up to 400 [C] and pressures up to 25 [MPa].
Fiber optic sensor (FOS) system used to capture distributed temperature measurements with spatial resolution ranging from 1.25 [mm] to 5 [mm].
Critical Heat Flux (CHF) events detected, and trigger shut down of experiment to protect heat rod and testing facility.
Health Monitoring via Distributed Temperature Measurements
Figure: (a) High pressure testing loop. (b) Experimental temperature profile [C] mapped on Position [m] vs Time [s] plot. White spots represent no data. (c) Temperature [C] vs Position plot [m] of CHF event.

## Slide 30: Distributed Strain Sensing
LUNA ODiSI-B

## Slide 31
Conduct research on embedding fiber optical sensors (FOS) onto the surface of tubes for monitoring internal pressure.
FOS are wrapped and bonded around tubes in a helical orientation or in a helical trench. The helical configuration allows for the optical strain sensor to be decomposed into its axial and circumferential components.
Develop a methodology for collecting strain data in high-temperature and high-pressure environments. The strain data needs to be compensated for temperature influence.
Health Monitoring via Distributed Strain Measurements
Figure: Schematic of helical trench on a 3/8 [in] O.D. tube w/ 22.5 [in] length.

## Slide 32: Experimental Setup and Correlation Methodology for the ODiSI-B
A single-mode fiber from Corning (SMF 28-Ultra) was wrapped in the helix groove and adhered with Ceramabond 668, a high-temperature ceramic adhesive.
Luna ODiSI-B was used to interrogate the SMF 28-Ultra fiber.
A hydraulic pump was used to pressurize the 3/8-inch SS tube up to 10,000± 100 [psi].
Frequency shift response to 1,000± 100 [psi] stages in internal pressure up to 10,000± 100 [psi] were measured.
Figure: Experimental setup (top). Frequency shift response to ten 1000± 100 [psi] stages(bottom).

## Slide 33: Luna ODiSI-B Correlation for Hoop Strain
Internal pressure was converted to hoop strain using pressurized vessel equations.
Hoop strain was plotted against frequency shift.
MS Excel was used to find a linear fit for a correlation.
New correlation is ε=-59.60υ where υ is the frequency shift values.
Figure: Hoop strain plotted against Frequency shift [dB]. MS Excel was used to find a linear fit.

## Slide 34: Thank you